# Supplementary Material

Machine Learning-Based Predictive Autofocus for High-Resolution OpenFlexure Microscopy Using Representative Defocus Ranges

K. Kathure, A. Kiroe, C. Ominde, D. M. Memeu

This document accompanies the main manuscript and repository. It provides (S1) extended methodological detail, (S2) full numerical data tables underlying Figures 1–6 of the manuscript (reported only as bar charts in the main text), and (S3) supplementary notes on scope and limitations.

## S1. Extended Methods

### S1.1 Engineered feature set

Six features were derived from a pair of images acquired at the current stage position and at a fixed offset of 200 motor steps (10 µm) along the Z-axis:

| S/NO | Feature | Definition |
| --- | --- | --- |
| 1 | Z_prev | Current Z-position of the stage (absolute positional context) |
| 2 | variance_ratio | VoL(true) / VoL(predicted) |
| 3 | variance_sq | VoL(true)² |
| 4 | direction | Sign of intended stage movement (+1 / −1) |
| 5 | is_improving | Binary flag: 1 if sharpness increased relative to the previous position, else 0 |
| 6 | absolute_variance | ¦VoL(true) − VoL(predicted)¦ |

Raw VoL values of the true/predicted images were used only to compute these six derived quantities and were not themselves passed to the models as separate inputs.

Variance of Laplacian (VoL):

### S1.2 Hyperparameter search spaces (grid search, 5-fold CV)

SVR (RBF kernel):

| Parameter | Search space |
| --- | --- |
| C | 0.1, 1, 10, 100, 1000 |
| γ | 0.001, 0.01, 0.1, 1, 10 |
| ε | 0.01, 0.1, 1, 10 |

Random Forest:

| Parameter | Search space |
| --- | --- |
| Number of trees (T) | 10, 50, 100, 200 |
| Max depth | None, 10, 20, 30, 50 |
| Min samples per split | 2, 5, 10, 20 |

Selection criterion: minimum validation-fold MSE. Features were standardized (zero mean, unit variance) for SVR only; Random Forest used raw features (tree splits are invariant to monotonic transforms).

### S1.3 Dataset partitioning

| Split | Samples |
| --- | --- |
| Total dataset | 1,040 images (8 fields of view) |
| Training (80%) | 832 |
| Test (20%, held out) | 208 |

Ground truth established via 4-stage hierarchical search: ultra-coarse (1000 µm range, 50 µm step) → coarse (200 µm range, 10 µm step) → fine (20 µm range, 2 µm step) → verification (±1 µm).

### S1.4 Hardware/optical configuration

| Component | Specification |
| --- | --- |
| Compute | Raspberry Pi 4 Model B |
| Camera | Raspberry Pi Camera Module V2, Sony IMX219, 8 MP, 1.12 µm pixel size |
| Acquisition resolution | 640 × 480 px |
| Objective (blood smears) | 100× oil immersion, NA 1.25 (Olympus UPlanFLN) |
| Objective (cancer tissue) | 40× dry, NA 0.65, FoV 320 × 240 µm |
| Stage resolution | 0.05 µm / motor step (calibrated) |
| Illumination | White LED, Köhler illumination, brightfield transmission |

## S2. Full Numerical Results (underlying Figures 1–6)

The main manuscript reports these results as grouped bar charts (Figs. 1–6). The tables below give the exact plotted values for reproducibility and reuse, grouped by sample type and initial defocus (motor steps). All values are means across n = 7–10 trials per condition; error bars (SD) are shown in the corresponding figures, not reproduced here.

### Table S1 — Final Variance of Laplacian (VoL, a.u.) — Fig. 1

| Sample | Defocus (steps) | SVR | Random Forest | Custom AF | Inbuilt AF |
| --- | --- | --- | --- | --- | --- |
| Stained | 500 | 115 | 98 | 110 | 108 |
| Stained | 1000 | 238 | 239 | 223 | 216 |
| Stained | 1500 | 70 | 69 | 109 | 49 |
| Unstained | 500 | 29.55 | 38.00 | 38.14 | 34.08 |
| Unstained | 1000 | 38.19 | 39.15 | 42.88 | 42.44 |
| Unstained | 1500 | 41.67 | 44.26 | 53.78 | 11.36 |
| Cancer tissue | 500 | 11.52 | 12.46 | 13.07 | 12.20 |
| Cancer tissue | 1000 | 10.76 | 15.68 | 26.74 | 9.55 |
| Cancer tissue | 1500 | 12.44 | 14.41 | 25.88 | 8.67 |

### Table S2 — Autofocus Execution Time (s) — Fig. 2

| Sample | Defocus (steps) | SVR | Random Forest | Custom AF | Inbuilt AF |
| --- | --- | --- | --- | --- | --- |
| Stained | 500 | 2.66 | 3.13 | 57.19 | 6.37 |
| Stained | 1000 | 2.80 | 2.12 | 66.08 | 6.49 |
| Stained | 1500 | 4.70 | 4.36 | 88.31 | 6.18 |
| Unstained | 500 | 3.36 | 3.56 | 118.08 | 7.40 |
| Unstained | 1000 | 4.25 | 4.30 | 68.09 | 5.78 |
| Unstained | 1500 | 4.30 | 4.53 | 70.51 | 5.74 |
| Cancer tissue | 500 | 4.20 | 4.18 | 47.08 | 6.06 |
| Cancer tissue | 1000 | 2.35 | 2.85 | 29.28 | 7.21 |
| Cancer tissue | 1500 | 3.47 | 3.86 | 32.91 | 7.16 |

### Table S3 — Coefficient of Variation of Final Z-Position (%) — Fig. 3

| Sample | Defocus (steps) | SVR | Random Forest | Custom AF | Inbuilt AF |
| --- | --- | --- | --- | --- | --- |
| Stained | 500 | 0.57 | 0.56 | 0.00 | 0.74 |
| Stained | 1000 | 0.02 | 0.00 | 0.80 | 0.26 |
| Stained | 1500 | 1.34 | 1.14 | 0.00 | 1.23 |
| Unstained | 500 | 0.00 | 0.41 | 2.81 | 5.70 |
| Unstained | 1000 | 0.09 | 1.34 | 0.53 | 0.37 |
| Unstained | 1500 | 0.09 | 0.02 | 0.72 | 0.31 |
| Cancer tissue | 500 | 0.11 | 0.83 | 1.77 | 9.89 |
| Cancer tissue | 1000 | 0.16 | 0.76 | 1.85 | 0.39 |
| Cancer tissue | 1500 | 0.62 | 0.34 | 0.64 | 0.32 |

### Table S4 — Relative Error to Reference Focal Plane (%) — Fig. 4

| Sample | Defocus (steps) | SVR | Random Forest | Custom AF | Inbuilt AF |
| --- | --- | --- | --- | --- | --- |
| Stained | 500 | 0.64 | 0.75 | 0.00 | 4.27 |
| Stained | 1000 | 2.25 | 2.26 | 0.73 | 3.79 |
| Stained | 1500 | 20.19 | 21.07 | 0.00 | 18.14 |
| Unstained | 500 | 2.12 | 1.33 | 1.87 | 8.42 |
| Unstained | 1000 | 3.41 | 2.78 | 0.49 | 5.12 |
| Unstained | 1500 | 5.72 | 5.78 | 0.56 | 5.34 |
| Cancer tissue | 500 | 3.87 | 5.39 | 1.55 | 16.15 |
| Cancer tissue | 1000 | 5.48 | 4.86 | 1.60 | 12.27 |
| Cancer tissue | 1500 | 7.81 | 7.54 | 0.44 | 0.97 |

### Table S5 — CPU Utilization (%) — Fig. 5

| Sample | Defocus (steps) | SVR | Random Forest | Custom AF | Inbuilt AF |
| --- | --- | --- | --- | --- | --- |
| Stained | 500 | 26.20 | 19.10 | 20.90 | 2.20 |
| Stained | 1000 | 51.90 | 51.60 | 45.90 | 1.90 |
| Stained | 1500 | 21.80 | 19.10 | 20.70 | 1.70 |
| Unstained | 500 | 20.98 | 28.19 | 27.33 | 1.64 |
| Unstained | 1000 | 22.93 | 22.95 | 27.31 | 2.42 |
| Unstained | 1500 | 23.36 | 23.90 | 20.89 | 2.52 |
| Cancer tissue | 500 | 35.63 | 34.29 | 34.56 | 3.79 |
| Cancer tissue | 1000 | 26.93 | 34.01 | 32.14 | 7.39 |
| Cancer tissue | 1500 | 31.06 | 36.43 | 29.40 | 8.62 |

### Table S6 — RAM Utilization (MB) — Fig. 6

| Sample | Defocus (steps) | SVR | Random Forest | Custom AF | Inbuilt AF |
| --- | --- | --- | --- | --- | --- |
| Stained | 500 | 0.10 | 0.20 | 1.20 | 2.70 |
| Stained | 1000 | 0.30 | 0.10 | 1.50 | 1.90 |
| Stained | 1500 | 0.20 | 0.40 | 0.50 | 2.00 |
| Unstained | 500 | 2.71 | 3.89 | 1.42 | 2.61 |
| Unstained | 1000 | 0.13 | 0.07 | 1.17 | 1.89 |
| Unstained | 1500 | 0.07 | 0.13 | 0.47 | 1.08 |
| Cancer tissue | 500 | 3.10 | 4.72 | 1.97 | 2.57 |
| Cancer tissue | 1000 | 3.42 | 3.00 | 1.78 | 2.51 |
| Cancer tissue | 1500 | 1.89 | 1.55 | 1.90 | 4.73 |

### Table S7 — Model Training and Cross-Validation Summary

| Model | Test MAE (steps) | Test MAE (µm) | Test RMSE (steps) | Test RMSE (µm) | Test R² | 5-fold CV mean R² |
| --- | --- | --- | --- | --- | --- | --- |
| Random Forest | 137.44 | 6.87 | 228.16 | 11.41 | 0.9814 | 0.9831 |
| SVR (RBF) | 224.87 | 11.24 | 291.78 | 14.58 | 0.9696 | 0.9656 |

Conversion assumes a calibrated displacement of 0.05 µm per motor step.

## S3. Supplementary Notes

### S3.1 Scope and reproducibility caveats

- All values in Tables S1–S6 were digitized directly from the manuscript’s own bar-chart figures (data labels shown on the plots); no values were estimated or interpolated.

- Sample sizes per condition were n = 7–10 trials; the tables above report means only. Refer to the manuscript’s figures for standard-deviation error bars.

- CV and relative-error values of exactly 0.00% for Custom AF reflect its role as the exhaustive-search ground-truth method in several conditions, not literal zero-variance measurement.

### S3.2 Known limitations (carried from manuscript )

- Each sample type was represented by a comparatively small number of fields of view and trials; broader staining protocols, tissue types, and cell densities remain to be tested.

- The accuracy degradation observed at 1500 motor steps (stained sample) indicates the effective operating range of the trained regression models is bounded below this offset; extending it will likely require additional training data at larger defocus magnitudes or a hybrid predictive + brief confirmatory-search strategy.

- Resource utilization (CPU/RAM) was measured on a single embedded hardware configuration (Raspberry Pi 4); results may differ on other single-board computers.

- Generalization findings on unstained and tissue-derived specimens would benefit from replication across a larger, more diverse specimen panel.

### S3.3 Data and code availability

The image dataset (1,040 labelled images, 8 fields of view) and trained model weights are available from the corresponding author upon reasonable request. Code for feature extraction, model training, and the live autofocus runtime is provided in this repository (see main README.md).

## Citation

If you use this supplementary data, please cite the main manuscript:

K. Kathure, A. Kiroe, C. Ominde, D. M. Memeu. "Machine Learning-Based Predictive
Autofocus for High-Resolution OpenFlexure Microscopy Using Representative Defocus
Ranges." Submitted to SLAS Technology, 2026.